## Instructions
SUPPLEMENTARY MATERIAL 4. Uncertainty calculation for the mass
concentration of clozapine in serum using EQAS to study the bias
UNCERTAINTY ESTIMATION SPREADSHEETS
Version 1 (Raúl Rigo-Bonnin, (0034) 93 260 75 43, raulr@bellvitgehospital.cat)
IMPORTANT NOTE
These spreadsheets have been developed with all care and attention. It does not come with a guarantee for being error-free in the workings or in the interpretation of the uncertainty calculation.
Please bring any errors or questions to the attention of Raúl Rigo-Bonnin
Phone: (0034) 93 260 75 43, raulr@bellvitgehospital.cat
Uncertainty estimation
The aim is to facilitate clinical laboratories the estimation of measurement uncertainty. This Supplementary material contains five different spreadsheeds. The first three spreadsheets allow to estimate the main measurement uncertainty sources (ucal, uRw and ub). The fourth have been generated to calculate the combined standard uncertainty as well as the expanded uncertainty, and to know if the values obtained accomplish with the measurement uncertainty requirements previously selected by the laboratory. The fifth describe how to specify the measurand.
Calculation Process
IMPORTANT NOTE: The process is OPERATOR DEPENDENT. Skill and attention to detail is important for
producing valid results. If in doubt seek expert advice and/or another operator for validation.
I am happy to be consulted for advice.
Data Entry
ENTRIES INTO THE SPREADSHEETS MUST ONLY BE MADE IN BLUE BOXES
1. | In the spreadshet "ucal", add the biological quantity, measuring system and measurement units into their respective boxes.
2. | In the spreadsheet "ucal", add the assigned values of the end-user calibrator, the expanded uncertainties (in units) provided by the manufacturer, and the coverage factors used by the manufacturer into their respective boxes.
3. | In the spreadsheet "uRw", add the data obtained by your laboratory. That is, for each IQC level, add the IQC mean values, the standard deviations, and the number of IQC values into their respective boxes.
4. | In the spreadsheet "ub", for every EQAS participation, add the measured values obtained by laboratory and the conventional values assigned by the EQAS manufacturer, into their respective boxes. Also, add the uncertainties associated with the reference value facilitated by the EQAS (in units) in their corresponding boxes. If these uncertainties are not facilitated, add the robust peer-group standard deviations and the number of peer-group laboratories in their corresponding boxes.
NOTE: It is important that if the boxes corresponding to the uncertainty associated with the reference value facilitated by the EQAS manufacturer are not used, they must not contain any value (that is, they must remain empty).
5. | In the spreadsheet "COMBINED & EXPANDED UNCERTAINTY", add the maximum permissible percent relative root mean square of measurement error obtained from RiliBÄK (if it exists) or, failing that, add the maximum allowable coefficient of variation (in %) and the maximum allowable percent relative bias (in %) in their corresponding boxes.
NOTE: It is important that if the box corresponding to the RiliBAK requirements is not used, it must not contain any value (that is, it must remains empty).

## ucal
Biological quantity: | Srm--Clozapine; subst.c.(USP; Acquity® UPLC®-TQD®)
Measuring system: | Acquity® UPLC®-TQD®
Measured values studied (IQC mean values): | 618 | 1689
Measurement unit: | nmo/L
Uncertainty related to the assigned value of the end-user calibration material
Measured value studied (x): | 618
Assigned value of the end-user calibrator (xcal): | 750
Expanded uncertainty (in units) provided by the manufacturer (Ucal): | 32.4972
Coverage factor used by the manufacturer (k): | 2
ucal = | 16.2486
Measured value studied (x): | 1689
Assigned value of the end-user calibrator (xcal): | 1500
Expanded uncertainty (in units) provided by the manufacturer (Ucal): | 57.375
Coverage factor used by the manufacturer (k): | 2
ucal = | 28.6875

## uRw
Biological quantity: | Srm--Clozapine; subst.c.(USP; Acquity® UPLC®-TQD®)
Measuring system 1: | Acquity® UPLC®-TQD®
Measured values studied for the estimation of measurement uncertainty (IQC mean values): | 618 | 1689
Measurement unit: | nmo/L
Uncertainty related to the long-term intermediate imprecision
Level 1
Measured value studied (x): | 618
IQC mean value obtained for the measuring system (`x ): | 618
Standard deviation obtained for the measuring system (s): | 28.0565
Number of IQC values processed for the measuring system (n): | 131
uRw = | 28.0565
Level 2
Measured value studied (x): | 1689
IQC mean value obtained for the measuring system (`x ): | 1689
Standard deviation obtained for the measuring system (s): | 36.8283
Number of IQC values processed for the measuring system (n): | 131
uRw = | 36.8283

## ub
Biological quantity: | Srm--Clozapine; subst.c.(USP; Acquity® UPLC®-TQD®)
Measuring system: | Acquity® UPLC®-TQD®
Measurement unit: | nmo/L
Uncertainty related to the bias using EQAS
EQAS participation (ni): | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Measured value obtained for the EQAS participation i (xi): | 408 | 291 | 216 | 918 | 516 | 589 | 2533 | 1750 | 715 | 575
Reference value (conventional value) assigned by the EQAS manufacturer in the EQAS participation i (mi): | 350 | 250 | 200 | 750 | 475 | 550 | 2225 | 1500 | 650 | 425
Measurement error obtained in the laboratory for the EQAS participation i (ei): | 58 | 41 | 16 | 168 | 41 | 39 | 308 | 250 | 65 | 150
Uncertainty associated with the reference value facilitated by the EQAS manufacturer for the EQAS participation i (umi ): | [6 empty] | 28.06 | 14.32 | 6.78 | 3.39
Robust peer-group standard deviation for the EQAS participation i (sLabs i): | 31.03 | 23.892 | 17.737 | 66.63 | 35.592 | 34.33
Number of peer-group laboratories included in the EQAS participation i(qi): | 92 | 88 | 92 | 90 | 93 | 89
Estimated uncertainty associated with the reference value for the EQAS participation i (um1): | 4.0439 | 3.1836 | 2.3115 | 8.7793 | 4.6134 | 4.5487 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Uncertainty associated with the reference value for the EQAS participation i (um1): | 4.0439 | 3.1836 | 2.3115 | 8.7793 | 4.6134 | 4.5487 | 28.06 | 14.32 | 6.78 | 3.39
b | 113.6
ub | 96.4195
Is the bias significant ? (½b½> 2 · ub): | NO

## COMBINED & EXPANDED UNCERTAINTY
Biological quantity: | Srm--Clozapine; subst.c.(USP; Acquity® UPLC®-TQD®)
Measuring system: | Acquity® UPLC®-TQD®
Measured values studied for the estimation of measurement uncertainty (IQC values): | 618 | 1689
Measurement unit: | nmo/L
Maximum permissible percent relative root mean square of measurement error (%Drel(max)) obtained from RiliBÄK (if applicable):
Maximum allowable coefficient of variation (CVmax): | 15 | 15
Maximum allowable percent relative bias (%brel(max)): | 15 | 15
Maximum allowable percent relative expanded uncertainty (%Urel(max)): | 21.2132 | 21.2132
Maximum allowable expanded uncertainty (Umax): | 131.0976 | 358.291
Combined standard uncertainty
Measured value studied (x): | 618 | 1689
Uncertainty related to the assigned value of the end-user calibrator (ucal): | 16.2486 | 28.6875
Uncertainty related to the long-term intermediate imprecision (uRw): | 28.0565 | 36.8283
Bias (b): | 0 | 0
uc = | 32.422 | 46.6829
%urel(c) = | 5.2463 | 2.7639
Expanded uncertainty
Measured value studied (x): | 618 | 1689
Combined standard uncertainty (uc): | 32.422 | 46.6829
Percent relative combined standard uncertainty (%rel(c)): | 5.2463 | 2.7639
Coverage factor (k): | 2 | 2
Measured value studied (x) = | 618 | 1689
U= | 64.8439 | 93.3659
%Urel = | 10.4925 | 5.5279
Is the expanded uncertainty obtained acceptable ? (U£ Umax): | YES | YES
Is the percent relative expanded uncertainty obtained acceptable ? (%Urel £ %Urel(max)): | YES | YES

## Specification of Measurand
The measurand can be defined as:
The substance concentration (in nmol/L) of clozapine in serum measured according to an in-house measurement procedure
using an Acquity® UPLC®-TQD® measuring system (Waters, Milford, MA, USA).
The results are traceable to assigned values of calibration materials prepared in HUB's lab using a clozapine CRM from
United States Phamacopeia (USP)
Furthermore, the quantity can also be described using the IUPAC-IFCC nomenclature as:
Srm¾Clozapine; subst.c.(USP; Acquity® UPLC®-TQD®)